Cryosectioning Tissue

- Pre-Cryosectioning Notes

  - You will need the following materials:

    - OCT tissue block (in cryomold)

    - OCT compound

    - Cryostat chucks

    - Microscope slides

    - Razor blades

    - Aluminum foil sections

  - Make sure the cryostat temperature is set to -20°C.

  - Keep the cryostat lid closed as much as possible for the entirety of this protocol. This will allow the internal temperature of the cryostat to remain around -20°C, reducing the chances of OCT liquifying and tissue freeze-thawing.

  - The blade guard should be in use when not sectioning.

  - OCT tissue blocks must remain either on dry ice or in the cryostat hood for the entirety of this protocol.

  - Make sure at least one chuck is inside the cryostat.

  - Label all microscope slides with hashcode, specimen number, and heart region.

- Protocol

- Place OCT tissue blocks in the cryostat and allow to equilibrate for 30 mins.

- Determine which face of the OCT tissue block you will be sectioning from.

  - If one face of the OCT tissue block has been sectioned from previously, section from the opposite face this iteration of the protocol. If both faces of the OCT tissue block have been sectioned from previously, section from the face that appears the most uniform this iteration of the protocol.

- Apply a dime-sized amount of OCT compound to the cryostat chuck, then place the OCT tissue block onto the chuck in the orientation with the chosen sectioning face outward.

- Allow the OCT compound to freeze to the chuck and OCT tissue block for ~10 seconds, then mount the chuck

- Allow time for the OCT to freeze the OCT tissue block to the chuck (~10 sec), then mount the chuck to the platform with one edge of the OCT tissue block parallel to the blade.

- Move the platform in the X-direction, towards the blade, using the corresponding buttons on the digital screen.

  - Stop moving the platform ~1 inch before the OCT tissue block and blade are overlapping in the Y-direction.

- Set the section thickness to 50 µm on the digital screen, then use the side handle/wheel to move the OCT tissue block closer to the blade and make ~2-3 sections. Discard these sections in the bin below the blade, and change the section thickness to 10 µm.

- Position the glass roll protector over the blade, and make a section.

- Lift up the glass roll protector and gently hover the corresponding room-temperature microscope slide above the section.

  - The section should “rise up” and adhere to the microscope slide.

- Repeat steps 8 and 9 for each section.

  - If “poor quality” sections are observed, refer to Troubleshooting.

- Once finished with an OCT tissue block, remove it from the chuck, re-wrap if with labeled aluminum foil, and place it on dry ice (if sectioning from another block) or in the -80°C freezer.

- Store all finished slides in -80°C freezer.

| Troubleshooting |  |  |
| --- | --- | --- |
| Problem | Possible Cause | Solution |
| Jagged cuts/tears in sections | Blade is chipped or dull | Change blade |
|  | Air pockets in OCT tissue block | Freeze more OCT on top of OCT tissue block |
|  | OCT tissue block is too cold | Allow additional time for OCT tissue block to equilibrate to cryostat temperature. |
| Sections “rolling up” | Sectioning too quickly | Slow down blade movement |
| Blade “jumping” over OCT tissue block | OCT tissue block is too cold. | Allow additional time for OCT tissue block to equilibrate to cryostat temperature. |
|  | Sectioning too quickly or too slowly. | Slow down or speed up blade movement. |