## Calibration Summary
Pipette | Leak test (pass/fail) | 100% Volume Tests (pass/fail) | 10% Volume Tests (pass/fail) | Calibration Needed? (yes/no)
10mL (1) | pass | PASS | PASS | NO
10mL (2) | pass | PASS | PASS | NO
10mL (3) | pass | PASS | PASS | NO
5mL (1) | pass | PASS | PASS | NO
5mL (2) | pass | PASS | PASS | NO
1000μL (1) | pass | PASS | FAIL | YES
1000μL (2) | ? | PASS | PASS | YES
1000μL (3) | ? | PASS | PASS | YES
1000μL (4) | pass | PASS | PASS | NO
1000μL (5) | pass | PASS | PASS | NO
1000μL (6) | pass | PASS | PASS | NO
300μL (1) | pass | PASS | PASS | NO
300μL (2) | pass | PASS | PASS | NO
300μL (3) | pass | PASS | FAIL | YES
100μL (1) | pass | PASS | FAIL | YES
100μL (2) | pass | PASS | FAIL | YES
100μL (3) | pass | PASS | PASS | NO
100μL (4) | ? | PASS | FAIL | YES
10μL (1) | pass | PASS | FAIL | YES
10μL (2) | pass | PASS | FAIL | YES
10μL (3) | pass | PASS | FAIL | YES
10μL (4) | pass | PASS | FAIL | YES
10μL (5) | pass | PASS | FAIL | YES
2.5μL (1) | pass | FAIL | FAIL | YES
2.5μL (2) | pass | FAIL | FAIL | YES
2.5μL (3) | pass | FAIL | FAIL | YES
2μL (1) | pass | FAIL | FAIL | YES
Systematic Error must be within ±0.6%, or pipette fails calibration

## 100% Volume Calculations
Pipette | Leak test (pass/fail) | Micrograms diH2O @ 100% volume |  |  |  | Mean Volume | Repeats | Systematic Error @ 100% Volume |  |  |  | Mean Sys. Error | Random Error @ 100% Volume |  |  |  | Mean Rand. Error | Coeefficient of Variation in % | Pass/Fail Calibration
10mL (1) | pass | 9.9611 | 9.9586 | 9.9735 | 10.0811 | 9.9936 | 4 | 0.32602 | 0.3512 | 0.20128 | -0.86821 | 0.0025742 | 0.0010546 | 0.0012233 | 0.00040301 | 0.0076606 | 0.058713 | 0.5875 | PASS
10mL (2) | pass | 9.98 | 9.9803 | 9.9631 | 9.9917 | 9.9788 | 4 | -0.012275 | -0.01528 | 0.15733 | -0.12936 | 0.00010463 | 1.5006e-06 | 2.3256e-06 | 0.00024571 | 0.00016706 | 0.011784 | 0.11809 | PASS
10mL (3) | pass | 9.9648 | 9.9674 | 9.9505 | 9.9539 | 9.9592 | 4 | -0.0567 | -0.08277 | 0.08693 | 0.052743 | 5.101e-05 | 3.1922e-05 | 6.8063e-05 | 7.4823e-05 | 2.7563e-05 | 0.0082132 | 0.082469 | PASS
5mL (1) | pass | 5.0172 | 5.0177 | 5.0361 | 5.0348 | 5.0265 | 4 | 0.18437 | 0.17438 | -0.19162 | -0.16585 | 0.00032156 | 8.5563e-05 | 7.6563e-05 | 9.3122e-05 | 6.9723e-05 | 0.010408 | 0.20706 | PASS
5mL (2) | pass | 4.9909 | 4.9979 | 5.0103 | 5.0144 | 5.0034 | 4 | 0.24995 | 0.10955 | -0.13822 | -0.21987 | 0.00035472 | 0.00015563 | 2.9976e-05 | 4.7956e-05 | 0.00012155 | 0.01088 | 0.21745 | PASS
1000μL (1) | pass | 1.0008 | 0.9993 | 1.003 | 0.975 | 0.99453 | 4 | -0.627 | -0.47783 | -0.84497 | 2.0026 | 0.013192 | 3.9376e-05 | 2.2801e-05 | 7.1826e-05 | 0.00038123 | 0.013105 | 1.3177 | PASS
1000μL (2) | ? | 0.965 | 0.9888 | 0.9826 | 0.9962 | 0.98315 | 4 | 1.8808 | -0.5714 | 0.055974 | -1.31 | 0.013856 | 0.00032942 | 3.1923e-05 | 3.025e-07 | 0.0001703 | 0.013316 | 1.3544 | PASS
1000μL (3) | ? | 1.0083 | 1.001 | 0.9994 | 1.004 | 1.0032 | 4 | -0.50828 | 0.21728 | 0.37773 | -0.082171 | 0.0011392 | 2.6266e-05 | 4.7306e-06 | 1.4251e-05 | 6.8062e-07 | 0.0039127 | 0.39003 | PASS
1000μL (4) | pass | 1.0046 | 1.0079 | 1.0094 | 1.0062 | 1.007 | 4 | 0.24139 | -0.086814 | -0.23529 | 0.081992 | 0.0003197 | 5.8806e-06 | 7.6562e-07 | 5.6406e-06 | 6.8062e-07 | 0.0020791 | 0.20646 | PASS
1000μL (5) | pass | 0.9773 | 0.9816 | 0.9974 | 0.983 | 0.98482 | 4 | 0.76998 | 0.32855 | -1.2608 | 0.18566 | 0.0058505 | 5.6626e-05 | 1.0401e-05 | 0.00015813 | 3.3306e-06 | 0.0087271 | 0.88616 | PASS
1000μL (6) | pass | 0.9638 | 1.0071 | 1.0162 | 0.9993 | 0.9966 | 4 | 3.4032 | -1.0426 | -1.9288 | -0.27019 | 0.040414 | 0.0010758 | 0.00011025 | 0.00038416 | 7.29e-06 | 0.022931 | 2.301 | PASS
300μL (1) | pass | 0.2956 | 0.3005 | 0.3032 | 0.2996 | 0.29973 | 4 | 1.3955 | -0.2579 | -1.1461 | 0.041722 | 0.0082944 | 1.7016e-05 | 6.0062e-07 | 1.2076e-05 | 1.5625e-08 | 0.0031468 | 1.0499 | PASS
300μL (2) | pass | 0.299 | 0.2988 | 0.2994 | 0.2985 | 0.29892 | 4 | -0.025084 | 0.041834 | -0.15865 | 0.14238 | 0.00011958 | 5.625e-09 | 1.5625e-08 | 2.2563e-07 | 1.8063e-07 | 0.00037749 | 0.12628 | PASS
300μL (3) | pass | 0.2987 | 0.2986 | 0.2985 | 0.301 | 0.2992 | 4 | 0.16739 | 0.20094 | 0.23451 | -0.59801 | 0.0012072 | 2.5e-07 | 3.6e-07 | 4.9e-07 | 3.24e-06 | 0.0012028 | 0.402 | PASS
100μL (1) | pass | 0.0991 | 0.0992 | 0.0997 | 0.1 | 0.0995 | 4 | 0.40363 | 0.30242 | -0.2006 | -0.5 | 0.0013626 | 1.6e-07 | 9e-08 | 4e-08 | 2.5e-07 | 0.00042426 | 0.4264 | PASS
100μL (2) | pass | 0.1004 | 0.1003 | 0.1002 | 0.1005 | 0.10035 | 4 | -0.049801 | 0.04985 | 0.1497 | -0.14925 | 0.00012413 | 2.5e-09 | 2.5e-09 | 2.25e-08 | 2.25e-08 | 0.0001291 | 0.12865 | PASS
100μL (3) | pass | 0.1 | 0.1005 | 0.1004 | 0.1007 | 0.1004 | 4 | 0.4 | -0.099502 | 5.529e-14 | -0.29791 | 0.00064573 | 1.6e-07 | 1e-08 | 3.0815e-33 | 9e-08 | 0.00029439 | 0.29322 | PASS
100μL (4) | ? | 0.0976 | 0.096 | 0.1002 | 0.0997 | 0.098375 | 4 | 0.79406 | 2.474 | -1.8214 | -1.329 | 0.029418 | 6.0063e-07 | 5.6406e-06 | 3.3306e-06 | 1.7556e-06 | 0.0019432 | 1.9752 | PASS
10μL (1) | pass | 0.0097 | 0.01 | 0.0099 | 0.0096 | 0.0098 | 4 | 1.0309 | -2 | -1.0101 | 2.0833 | 0.02604 | 1e-08 | 4e-08 | 1e-08 | 4e-08 | 0.00018257 | 1.863 | PASS
10μL (2) | pass | 0.0096 | 0.01 | 0.0095 | 0.0098 | 0.009725 | 4 | 1.3021 | -2.75 | 2.3684 | -0.76531 | 0.0388 | 1.5625e-08 | 7.5625e-08 | 5.0625e-08 | 5.625e-09 | 0.00022174 | 2.2801 | PASS
10μL (3) | pass | 0.0096 | 0.0101 | 0.0099 | 0.0098 | 0.00985 | 4 | 2.6042 | -2.4752 | -0.50505 | 0.5102 | 0.033518 | 6.25e-08 | 6.25e-08 | 2.5e-09 | 2.5e-09 | 0.00020817 | 2.1134 | PASS
10μL (4) | pass | 0.01 | 0.0102 | 0.0098 | 0.0099 | 0.009975 | 4 | -0.25 | -2.2059 | 1.7857 | 0.75758 | 0.021852 | 6.25e-10 | 5.0625e-08 | 3.0625e-08 | 5.625e-09 | 0.00017078 | 1.7121 | PASS
10μL (5) | pass | 0.0101 | 0.0102 | 0.01 | 0.0097 | 0.01 | 4 | -0.9901 | -1.9608 | 0 | 3.0928 | 0.035475 | 1e-08 | 4e-08 | 0 | 9e-08 | 0.00021602 | 2.1602 | PASS
2.5μL (1) | pass | 0.0025 | 0.0032 | 0.0026 | 0.0027 | 0.00275 | 4 | 10 | -14.0625 | 5.7692 | 1.8519 | 0.88965 | 6.25e-08 | 2.025e-07 | 2.25e-08 | 2.5e-09 | 0.00031091 | 11.3059 | FAIL
2.5μL (2) | pass | 0.0025 | 0.0024 | 0.0023 | 0.0027 | 0.002475 | 4 | -1 | 3.125 | 7.6087 | -8.3333 | 0.35009 | 6.25e-10 | 5.625e-09 | 3.0625e-08 | 5.0625e-08 | 0.00017078 | 6.9003 | FAIL
2.5μL (3) | pass | 0.0023 | 0.0027 | 0.0023 | 0.0025 | 0.00245 | 4 | 6.5217 | -9.2593 | 6.5217 | -2 | 0.44605 | 2.25e-08 | 6.25e-08 | 2.25e-08 | 2.5e-09 | 0.00019149 | 7.8157 | FAIL
2μL (1) | pass | 0.0019 | 0.0019 | 0.002 | 0.0024 | 0.00205 | 4 | 7.8947 | 7.8947 | 2.5 | -14.5833 | 0.92654 | 2.25e-08 | 2.25e-08 | 2.5e-09 | 1.225e-07 | 0.00023805 | 11.6121 | FAIL
Systematic Error must be within ±0.6%, or pipette fails calibration
Note on I1: Yackzan, Andrew T.: Systematic Error must be within ±0.6%, or pipette fails calibration

## 10% Volume Calculations
Pipette | Leak test (pass/fail) | Micrograms diH2O @ 10% volume |  |  |  | Mean Volume | Repeats | Systemic Error @ 10% Volume |  |  |  | Mean Sys. Error | Precision @ 10% Volume |  |  |  | Mean Rand. Error | Coeefficient of Variation in % | Pass/Fail Calibration
10mL (1) | pass | 1.0119 | 1.0094 | 1.0037 | 1.0159 | 1.0102 | 4 | -0.16553 | 0.081732 | 0.65009 | -0.55862 | 0.0019196 | 2.8056e-06 | 6.8063e-07 | 4.2576e-05 | 3.2206e-05 | 0.0051078 | 0.50561 | PASS
10mL (2) | pass | 1.0264 | 1.0359 | 1.0297 | 1.0201 | 1.028 | 4 | 0.15832 | -0.76021 | -0.16267 | 0.77688 | 0.0030819 | 2.6406e-06 | 6.2016e-05 | 2.8056e-06 | 6.2806e-05 | 0.0065896 | 0.64099 | PASS
10mL (3) | pass | 0.9978 | 1.0049 | 1.0099 | 1.0111 | 1.0059 | 4 | 0.81429 | 0.102 | -0.3936 | -0.51182 | 0.0027174 | 6.6016e-05 | 1.0506e-06 | 1.5801e-05 | 2.6781e-05 | 0.0060456 | 0.601 | PASS
5mL (1) | pass | 1.0039 | 1.0025 | 0.9813 | 0.9971 | 0.9962 | 4 | -0.76701 | -0.62843 | 1.5184 | -0.090262 | 0.0081737 | 5.929e-05 | 3.969e-05 | 0.00022201 | 8.1e-07 | 0.010357 | 1.0396 | PASS
5mL (2) | pass | 0.9997 | 1.0259 | 1.0145 | 1.0062 | 1.0116 | 4 | 1.1879 | -1.3963 | -0.28832 | 0.53419 | 0.0093475 | 0.00014102 | 0.00020521 | 8.5556e-06 | 2.8891e-05 | 0.011309 | 1.1179 | PASS
1000μL (1) | pass | 0.1048 | 0.104 | 0.1047 | 0.0952 | 0.10217 | 4 | -2.5048 | -1.7548 | -2.4117 | 7.3267 | 0.16386 | 6.8906e-06 | 3.3306e-06 | 6.3756e-06 | 4.8651e-05 | 0.0046636 | 4.5643 | FAIL
1000μL (2) | ? | 0.1049 | 0.1035 | 0.1033 | 0.1023 | 0.1035 | 4 | -1.3346 | 0 | 0.19361 | 1.173 | 0.0080067 | 1.96e-06 | 0 | 4e-08 | 1.44e-06 | 0.0010708 | 1.0346 | PASS
1000μL (3) | ? | 0.1019 | 0.1012 | 0.0967 | 0.1023 | 0.10053 | 4 | -1.3494 | -0.667 | 3.9555 | -1.7351 | 0.05102 | 1.8906e-06 | 4.5563e-07 | 1.4631e-05 | 3.1506e-06 | 0.0025902 | 2.5767 | PASS
1000μL (4) | pass | 0.1054 | 0.1032 | 0.1026 | 0.1031 | 0.10358 | 4 | -1.7315 | 0.36337 | 0.95029 | 0.46072 | 0.010721 | 3.3306e-06 | 1.4062e-07 | 9.5062e-07 | 2.2563e-07 | 0.0012447 | 1.2017 | PASS
1000μL (5) | pass | 0.1048 | 0.1041 | 0.1068 | 0.1036 | 0.10483 | 4 | 0.023855 | 0.69645 | -1.8493 | 1.1824 | 0.013371 | 6.25e-10 | 5.2562e-07 | 3.9006e-06 | 1.5006e-06 | 0.0014056 | 1.3409 | PASS
1000μL (6) | pass | 0.1041 | 0.1035 | 0.1077 | 0.1067 | 0.1055 | 4 | 1.3449 | 1.9324 | -2.0427 | -1.1246 | 0.027467 | 1.96e-06 | 4e-06 | 4.84e-06 | 1.44e-06 | 0.0020199 | 1.9146 | PASS
300μL (1) | pass | 0.0307 | 0.0295 | 0.0311 | 0.0309 | 0.03055 | 4 | -0.4886 | 3.5593 | -1.7685 | -1.1327 | 0.042387 | 2.25e-08 | 1.1025e-06 | 3.025e-07 | 1.225e-07 | 0.0007188 | 2.3528 | PASS
300μL (2) | pass | 0.0302 | 0.03 | 0.0301 | 0.0314 | 0.030425 | 4 | 0.74503 | 1.4167 | 1.0797 | -3.1051 | 0.034085 | 5.0625e-08 | 1.8063e-07 | 1.0562e-07 | 9.5063e-07 | 0.00065511 | 2.1532 | PASS
300μL (3) | pass | 0.0303 | 0.033 | 0.0296 | 0.0295 | 0.0306 | 4 | 0.9901 | -7.2727 | 3.3784 | 3.7288 | 0.20614 | 9e-08 | 5.76e-06 | 1e-06 | 1.21e-06 | 0.0016391 | 5.3566 | FAIL
100μL (1) | pass | 0.0096 | 0.0098 | 0.0103 | 0.0097 | 0.00985 | 4 | 2.6042 | 0.5102 | -4.3689 | 1.5464 | 0.072958 | 6.25e-08 | 2.5e-09 | 2.025e-07 | 2.25e-08 | 0.00031091 | 3.1565 | FAIL
100μL (2) | pass | 0.0104 | 0.0129 | 0.0098 | 0.01 | 0.010775 | 4 | 3.6058 | -16.4729 | 9.949 | 7.75 | 1.208 | 1.4063e-07 | 4.5156e-06 | 9.5062e-07 | 6.0062e-07 | 0.0014385 | 13.35 | FAIL
100μL (3) | pass | 0.01 | 0.0101 | 0.0101 | 0.0097 | 0.009975 | 4 | -0.25 | -1.2376 | -1.2376 | 2.8351 | 0.027451 | 6.25e-10 | 1.5625e-08 | 1.5625e-08 | 7.5625e-08 | 0.0001893 | 1.8977 | PASS
100μL (4) | ? | 0.0115 | 0.0107 | 0.0098 | 0.0095 | 0.010375 | 4 | -9.7826 | -3.0374 | 5.8673 | 9.2105 | 0.56447 | 1.2656e-06 | 1.0563e-07 | 3.3062e-07 | 7.6563e-07 | 0.00090692 | 8.7414 | FAIL
10μL (1) | pass | 0.0007 | 0.0012 | 0.0008 | 0.0007 | 0.00085 | 4 | 21.4286 | -29.1667 | 6.25 | 21.4286 | 4.9851 | 2.25e-08 | 1.225e-07 | 2.5e-09 | 2.25e-08 | 0.00023805 | 28.0056 | FAIL
10μL (2) | pass | 0.0009 | 0.0007 | 0.0008 | 0.0008 | 0.0008 | 4 | -11.1111 | 14.2857 | 8.6736e-13 | 8.6736e-13 | 0.79365 | 1e-08 | 1e-08 | 4.8148e-35 | 4.8148e-35 | 8.165e-05 | 10.2062 | FAIL
10μL (3) | pass | 0.001 | 0.001 | 0.0009 | 0.001 | 0.000975 | 4 | -2.5 | -2.5 | 8.3333 | -2.5 | 0.20833 | 6.25e-10 | 6.25e-10 | 5.625e-09 | 6.25e-10 | 5e-05 | 5.1282 | FAIL
10μL (4) | pass | 0.0015 | 0.0008 | 0.001 | 0.0011 | 0.0011 | 4 | -26.6667 | 37.5 | 10 | -6.3081e-13 | 5.2083 | 1.6e-07 | 9e-08 | 1e-08 | 4.8148e-35 | 0.00029439 | 26.7629 | FAIL
10μL (5) | pass | 0.0015 | 0.0013 | 0.0013 | 0.0004 | 0.001125 | 4 | -25 | -13.4615 | -13.4615 | 181.25 | 32.3317 | 1.4063e-07 | 3.0625e-08 | 3.0625e-08 | 5.2563e-07 | 0.00049244 | 43.7727 | FAIL
2.5μL (1) | pass | 0.0005 | 0.0001 | 0 | 0.0006 | 0.0003 | 4 | -40 | 200 | NaN | -50 | 36.6667 | 4e-08 | 4e-08 | 9e-08 | 9e-08 | 0.00029439 | 98.1307 | FAIL
2.5μL (2) | pass | 0 | 0.0002 | 0 | 0.0007 | 0.000225 | 4 | NaN | 12.5 | NaN | -67.8571 | -27.6786 | 5.0625e-08 | 6.25e-10 | 5.0625e-08 | 2.2563e-07 | 0.0003304 | 146.8461 | FAIL
2.5μL (3) | pass | 0.0003 | 0.0003 | 0.0002 | 0.0001 | 0.000225 | 4 | -25 | -25 | 12.5 | 125 | 21.875 | 5.625e-09 | 5.625e-09 | 6.25e-10 | 1.5625e-08 | 9.5743e-05 | 42.5523 | FAIL
2μL (1) | pass | 0.0003 | 0.0001 | 0 | 0 | 0.0001 | 4 | -66.6667 | -1.3878e-11 | NaN | NaN | -33.3333 | 4e-08 | 1.9259e-34 | 1e-08 | 1e-08 | 0.00014142 | 141.4214 | FAIL
Systematic Error must be within ±0.6%, or pipette fails calibration
Note on I1: Yackzan, Andrew T.: Systematic Error must be within ±0.6%, or pipette fails calibration
Note on C5: Yackzan, Andrew T.: 20% max volume (1mL)
Note on D5: Yackzan, Andrew T.: 20% max volume (1mL)
Note on E5: Yackzan, Andrew T.: 20% max volume (1mL)
Note on F5: Yackzan, Andrew T.: 20% max volume (1mL)
Note on G5: Yackzan, Andrew T.: 20% max volume (1mL)
Note on C6: Yackzan, Andrew T.: 20% max volume (1mL)
Note on D6: Yackzan, Andrew T.: 20% max volume (1mL)
Note on E6: Yackzan, Andrew T.: 20% max volume (1mL)
Note on F6: Yackzan, Andrew T.: 20% max volume (1mL)
Note on G6: Yackzan, Andrew T.: 20% max volume (1mL)